## README
Structure of the file:
This Excel file contains 4 sheets in addition to this Readme:
1) CTnumbers
2) PhantomInserts
3) TabulatedHumanTissues
4) ElementParameters
Content of each sheet:
"CTnumbers" is blank, but here the mean CT numbers extracted ROIs placed in each phantom inserts in the CT scans should be filled in. This Excel file is set up for the CIRS EPTN phantom. The listed phantom inserts need to match those listed in sheet "PhantomInserts".
"PhantomInserts" contains the elemental composition for the phantom inserts as provided by the phantom vendor. The mass densities for the phantom inserts should be filled in here. If the SPR values for the phantom inserts have been measured, these should also be filled in here.
"TabulatedHumanTissues" contains the elemental composition and mass densities for the tabulated human tissues, as given by Woodard and White, https://doi.org/10.1259/0007-1285-59-708-1209.
"ElementParameters" contains chemical information of the different elements.
Important notes:
1. The material information for the phantom inserts was current at the time of publication. Users should confirm that it reflects their own phantom, since phantom kits and/or materials could change.
2. CT numbers will vary based upon the scanner, scan protocols, phantom setup, etc.
User instructions:
1. Check that the inserts names match the inserts in your phantom. If another phantom is used, the names should be changed accordingly. The phantom inserts names should be updated both in the sheet "CTnumbers" and "PhantomInserts". And the order of the inserts should be the same in the two sheets.
2. Fill in the mean CT numbers measured for your phantom in the sheet "CTnumbers".
3. Fill in the mass density for your phantom inserts in the sheet "PhantomInserts", in column D.
4. If you have measured the SPR for your phantom inserts, the measured SPR values should be filled into the sheet "PhantomInserts", in column E.
5. Check that the elemental composition and mass density of the phantom inserts matches your phantom inserts, as specified on the datasheet received together with your phantom. Here batch specific data should be used. Especially the mass density might change. If another phantom model is used, all data should be changed.
Notes on use:
1. Make sure the phantom inserts are listed in the same order in the sheet "CTnumbers" and "PhantomInserts".
2. The order of the elements should be the same in the three sheets "CTnumbers", "PhantomInserts", and "ElementParameters".
3. If the names of the sheets, or the row headings in any of the sheets are changed, the Matlab and Python scripts also need to be changed accordingly.
Extra information:
More information on the CT scanning, SPR measurements, and HLUT fitting process can be found in the article: DOI: https://doi.org/10.1016/j.radonc.2023.109675

## CTnumbers
Insert name | CT number (Head) | CT number (Body) | CT number (Body periphery)
Lung inhale
Lung medium
Adipose
30% Gland + 70% Adipose
70% Gland + 30% Adipose
Gland
Water
Solid water
Soft Tissue Grey
Spinal Cord
Kidney
Prostate
Brain tissue
Spinal Disk
Trabecular Bone
30% Cortical Bone + 70% Trabecular Bone
70% Cortical Bone + 30% Trabecular Bone
Cortical Bone

## PhantomInserts
Tissue type | Tissue group | Insert name | Density (g/cm3) | SPR Measured | H | Li | Be | B | C | N | O | F | Na | Mg | Al | Si | P | S | Cl | K | Ca | Ti | Fe | Zn | I | Ba
Lung tissue | 1 | Lung inhale |  |  | 0.0859 | 0 | 0 | 0 | 0.6589 | 0.0352 | 0.1926 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0169 | 0 | 0.0101 | 0 | 0 | 0 | 0 | 0
| 1 | Lung medium |  |  | 0.0791 | 0 | 0 | 0 | 0.5478 | 0.0247 | 0.2571 | 0 | 0 | 0.0814 | 0 | 0 | 0 | 0 | 0.0095 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fat tissue | 2 | Adipose |  |  | 0.0997 | 0 | 0 | 0 | 0.7136 | 0.0179 | 0.1637 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0019 | 0 | 0.0032 | 0 | 0 | 0 | 0 | 0
| 2 | 30% Gland + 70% Adipose |  |  | 0.0967 | 0 | 0 | 0 | 0.7072 | 0.0196 | 0.1673 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0022 | 0 | 0.0062 | 0 | 0 | 0 | 0 | 0
| 2 | 70% Gland + 30% Adipose |  |  | 0.0952 | 0 | 0 | 0 | 0.698 | 0.019 | 0.1726 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0018 | 0 | 0.0126 | 0 | 0 | 0 | 0 | 0
| 2 | Gland |  |  | 0.0941 | 0 | 0 | 0 | 0.6913 | 0.0184 | 0.1766 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0014 | 0 | 0.0175 | 0 | 0 | 0 | 0 | 0
Soft tissue | 3 | Water |  |  | 0.1119 | 0 | 0 | 0 | 0 | 0 | 0.8881 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Solid water |  |  | 0.080022 | 0 | 0 | 0 | 0.67301 | 0.023888 | 0.19871 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0013627 | 0 | 0.023103 | 0 | 0 | 0 | 0 | 0
| 3 | Soft Tissue Grey |  |  | 0.0847 | 0 | 0 | 0 | 0.5744 | 0.0165 | 0.2459 | 0 | 0 | 0.0762 | 0 | 0 | 0 | 0 | 0.0019 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Spinal Cord |  |  | 0.0846 | 0 | 0 | 0 | 0.5621 | 0.0153 | 0.2478 | 0 | 0 | 0.0841 | 0 | 0 | 0.0012 | 0 | 0.0018 | 0 | 0.0025 | 0 | 0 | 0 | 0 | 0
| 3 | Kidney |  |  | 0.0852 | 0 | 0 | 0 | 0.6295 | 0.0197 | 0.2063 | 0 | 0 | 0.0445 | 0 | 0 | 0 | 0 | 0.0016 | 0 | 0.013 | 0 | 0 | 0 | 0 | 0
| 3 | Prostate |  |  | 0.0744 | 0 | 0 | 0 | 0.558 | 0.0224 | 0.259 | 0 | 0 | 0.0841 | 0 | 0 | 0 | 0 | 0.0021 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Brain tissue |  |  | 0.0816 | 0 | 0 | 0 | 0.536 | 0.0153 | 0.2649 | 0 | 0 | 0.0998 | 0 | 0 | 0 | 0 | 0.0019 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bone tissue | 4 | Spinal Disk |  |  | 0.0671 | 0 | 0 | 0 | 0.4576 | 0.0188 | 0.3106 | 0 | 0 | 0.1436 | 0 | 0 | 0 | 0 | 0.0021 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 4 | Trabecular Bone |  |  | 0.0746 | 0 | 0 | 0 | 0.5484 | 0.013 | 0.2528 | 0 | 0 | 0.0212 | 0 | 0 | 0 | 0 | 0.0011 | 0 | 0.0884 | 0 | 0 | 0 | 0 | 0
| 4 | 30% Cortical Bone + 70% Trabecular Bone |  |  | 0.0653 | 0 | 0 | 0 | 0.4985 | 0.0151 | 0.2376 | 0 | 0 | 0 | 0 | 0 | 0.0556 | 0 | 0.0007 | 0 | 0.1269 | 0 | 0 | 0 | 0 | 0
| 4 | 70% Cortical Bone + 30% Trabecular Bone |  |  | 0.0488 | 0 | 0 | 0 | 0.3692 | 0.0117 | 0.2913 | 0 | 0 | 0.0121 | 0 | 0 | 0.0819 | 0 | 0.0004 | 0 | 0.1844 | 0 | 0 | 0 | 0 | 0
| 4 | Cortical Bone |  |  | 0.033 | 0 | 0 | 0 | 0.2537 | 0.0091 | 0.3528 | 0 | 0 | 0.0336 | 0 | 0 | 0.0882 | 0 | 0.0003 | 0 | 0.2291 | 0 | 0 | 0 | 0 | 0

## TabulatedHumanTissues
Tissue type | Tissue group | Tissue name | Density (g/cm3) | H | Li | Be | B | C | N | O | F | Na | Mg | Al | Si | P | S | Cl | K | Ca | Ti | Fe | Zn | I | Ba
Lung tissue | 1 | Lung inflated | 0.26 | 0.103 | 0 | 0 | 0 | 0.105 | 0.031 | 0.749 | 0 | 0.002 | 0 | 0 | 0 | 0.002 | 0.003 | 0.003 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0
Fat tissue | 2 | Adipose tissue 3​ | 0.93 | 0.116 | 0 | 0 | 0 | 0.681 | 0.002 | 0.198 | 0 | 0.001 | 0 | 0 | 0 | 0 | 0.001 | 0.001 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 2 | Adipose tissue 2 | 0.95 | 0.114 | 0 | 0 | 0 | 0.598 | 0.007 | 0.278 | 0 | 0.001 | 0 | 0 | 0 | 0 | 0.001 | 0.001 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 2 | Adipose tissue 1 | 0.97 | 0.112 | 0 | 0 | 0 | 0.517 | 0.013 | 0.355 | 0 | 0.001 | 0 | 0 | 0 | 0 | 0.001 | 0.001 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 2 | Skeleton – yellow marrow​ | 0.98 | 0.115 | 0 | 0 | 0 | 0.644 | 0.007 | 0.231 | 0 | 0.001 | 0 | 0 | 0 | 0 | 0.001 | 0.001 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 2 | Mammary gland​ 2 | 1.02 | 0.106 | 0 | 0 | 0 | 0.332 | 0.03 | 0.527 | 0 | 0.001 | 0 | 0 | 0 | 0.001 | 0.002 | 0.001 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 2 | Skeleton – red marrow​ | 1.03 | 0.105 | 0 | 0 | 0 | 0.414 | 0.034 | 0.439 | 0 | 0 | 0 | 0 | 0 | 0.001 | 0.002 | 0.002 | 0.002 | 0 | 0 | 0.001 | 0 | 0 | 0
| 2 | Adrenal gland​ | 1.03 | 0.106 | 0 | 0 | 0 | 0.284 | 0.026 | 0.578 | 0 | 0 | 0 | 0 | 0 | 0.001 | 0.002 | 0.002 | 0.001 | 0 | 0 | 0 | 0 | 0 | 0
Soft tissue | 3 | Brain – Cerebrospinal Fluid | 1.01 | 0.111 | 0 | 0 | 0 | 0 | 0 | 0.88 | 0 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0.004 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Small intestine (wall)​ | 1.03 | 0.106 | 0 | 0 | 0 | 0.115 | 0.022 | 0.751 | 0 | 0.001 | 0 | 0 | 0 | 0.001 | 0.001 | 0.002 | 0.001 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Urinary bladder – filled | 1.03 | 0.108 | 0 | 0 | 0 | 0.035 | 0.015 | 0.83 | 0 | 0.003 | 0 | 0 | 0 | 0.001 | 0.001 | 0.005 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Gallbladder – bile​ | 1.03 | 0.108 | 0 | 0 | 0 | 0.061 | 0.001 | 0.822 | 0 | 0.004 | 0 | 0 | 0 | 0 | 0 | 0.004 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Pancreas​ | 1.04 | 0.106 | 0 | 0 | 0 | 0.169 | 0.022 | 0.694 | 0 | 0.002 | 0 | 0 | 0 | 0.002 | 0.001 | 0.002 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Brain – white matter​ | 1.04 | 0.106 | 0 | 0 | 0 | 0.194 | 0.025 | 0.661 | 0 | 0.002 | 0 | 0 | 0 | 0.004 | 0.002 | 0.003 | 0.003 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Prostate | 1.04 | 0.105 | 0 | 0 | 0 | 0.089 | 0.025 | 0.774 | 0 | 0.002 | 0 | 0 | 0 | 0.001 | 0.002 | 0 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Testis | 1.04 | 0.106 | 0 | 0 | 0 | 0.099 | 0.02 | 0.766 | 0 | 0.002 | 0 | 0 | 0 | 0.001 | 0.002 | 0.002 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Brain – grey matter | 1.04 | 0.107 | 0 | 0 | 0 | 0.095 | 0.018 | 0.767 | 0 | 0.002 | 0 | 0 | 0 | 0.003 | 0.002 | 0.003 | 0.003 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Stomach | 1.05 | 0.104 | 0 | 0 | 0 | 0.139 | 0.029 | 0.721 | 0 | 0.001 | 0 | 0 | 0 | 0.001 | 0.002 | 0.001 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Muscle – skeletal​ 2 | 1.05 | 0.102 | 0 | 0 | 0 | 0.143 | 0.034 | 0.71 | 0 | 0.001 | 0 | 0 | 0 | 0.002 | 0.003 | 0.001 | 0.004 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Heart​ 2 | 1.05 | 0.104 | 0 | 0 | 0 | 0.139 | 0.029 | 0.718 | 0 | 0.001 | 0 | 0 | 0 | 0.002 | 0.002 | 0.002 | 0.003 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Aorta | 1.05 | 0.099 | 0 | 0 | 0 | 0.147 | 0.042 | 0.698 | 0 | 0.002 | 0 | 0 | 0 | 0.004 | 0.003 | 0 | 0.001 | 0.004 | 0 | 0 | 0 | 0 | 0
| 3 | Kidney​ 2 | 1.05 | 0.103 | 0 | 0 | 0 | 0.132 | 0.03 | 0.724 | 0 | 0.002 | 0 | 0 | 0 | 0.002 | 0.002 | 0.002 | 0.002 | 0.001 | 0 | 0 | 0 | 0 | 0
| 3 | Ovary | 1.05 | 0.105 | 0 | 0 | 0 | 0.093 | 0.024 | 0.768 | 0 | 0.002 | 0 | 0 | 0 | 0.002 | 0.002 | 0.002 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Liver 2 | 1.06 | 0.102 | 0 | 0 | 0 | 0.139 | 0.03 | 0.716 | 0 | 0.002 | 0 | 0 | 0 | 0.003 | 0.003 | 0.002 | 0.003 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Spleen | 1.06 | 0.103 | 0 | 0 | 0 | 0.113 | 0.032 | 0.741 | 0 | 0.001 | 0 | 0 | 0 | 0.003 | 0.002 | 0.002 | 0.003 | 0 | 0 | 0 | 0 | 0 | 0
| 3 | Blood – whole​ | 1.06 | 0.102 | 0 | 0 | 0 | 0.11 | 0.033 | 0.745 | 0 | 0.001 | 0 | 0 | 0 | 0.001 | 0.002 | 0.003 | 0.002 | 0 | 0 | 0.001 | 0 | 0 | 0
Bone tissue | 4 | Skeleton – cartilage​ | 1.1 | 0.096 | 0 | 0 | 0 | 0.099 | 0.022 | 0.744 | 0 | 0.005 | 0 | 0 | 0 | 0.022 | 0.009 | 0.003 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 4 | Skeleton – spongiosa​ | 1.18 | 0.085 | 0 | 0 | 0 | 0.404 | 0.028 | 0.367 | 0 | 0.001 | 0.001 | 0 | 0 | 0.034 | 0.002 | 0.002 | 0.001 | 0.074 | 0 | 0.001 | 0 | 0 | 0
| 4 | Skeleton – cortical bone​ | 1.92 | 0.034 | 0 | 0 | 0 | 0.155 | 0.042 | 0.435 | 0 | 0.001 | 0.002 | 0 | 0 | 0.103 | 0.003 | 0 | 0 | 0.225 | 0 | 0 | 0 | 0 | 0

## ElementParameters
Element | Zi | Ai | Ii
H | 1 | 1.008 | 19.2
Li | 3 | 6.94 | 40
Be | 4 | 9.0122 | 63.7
B | 5 | 10.81 | 76
C | 6 | 12.011 | 81
N | 7 | 14.007 | 82
O | 8 | 15.999 | 106
F | 9 | 18.998 | 112
Na | 11 | 22.99 | 149
Mg | 12 | 24.305 | 156
Al | 13 | 26.982 | 166
Si | 14 | 28.085 | 173
P | 15 | 30.974 | 173
S | 16 | 32.06 | 180
Cl | 17 | 35.45 | 180
K | 19 | 39.098 | 190
Ca | 20 | 40.078 | 191
Ti | 22 | 47.867 | 233
Fe | 26 | 55.845 | 286
Zn | 30 | 65.38 | 330
I | 53 | 126.9 | 554.83
Ba | 56 | 137.33 | 491