## Description
CT scans
CT scans of a tissue-equivalent guide-compliant phantom (Gammex Advanced Electron Density Phantom, Sun Nuclear) and insufficiently tissue-equivalent phantom used in the Letter by Poludniowski and Zimmerman (CIRS 062M, Sun Nuclear) were acquired with a Somatom Confidence CT scanner (Siemens Healthineers, Forchheim, Germany) at 120 kVp and reconstructed with Qr40 kernel, SAFIRE level 3, and iBHC of bone. Following the consensus guide, the soft tissues were scanned separately, while the bone inserts were scanned one at a time, placed in the middle of the phantom. The scans of phantom configurations including bone inserts were only used to extract the CT number for the bone insert, while the surrounding soft tissue inserts were ignored. The CT numbers of the soft tissue inserts were extracted from the one phantom configuration only including soft tissues. For both phantoms, only the inner head-sized part was scanned.
Data extraction
The mean CT numbers were extracted from regions-of-interest (ROIs) covering 70% of the phantom diameter and several axial CT slices around the center of the phantom. The extracted CT numbers are listed in the sheet "PhantomInserts" for both phantoms. Here the batch-specific densities of the phantom inserts and their chemical compositions are also stated, as well as their stopping-power ratio (SPR) at 100 MeV as calculated with the Bethe equation.
CT number estimation
For each phantom separately, the measured CT numbers and the chemical compositions of the phantom inserts were used to calibrate two different CT number estimation methods, as described by Scheider et al., Phys Med Biol (1996), and by Ödén et al., Phys Med (2018). In the latter article several CT number estimation methods are described, the method used here is the one termed the Karolinska model.
Each of the resulting four CT number estimation methods were applied to the tabulated human tissues given by Woodard and White, Br J Radiol (1986) and listed in the consensus guide. These estimated CT numbers are listed in the sheet "TabulatedHumanTissues".
Hounsfield look-up table (HLUT) specification
To investigate the influence of the tissue equivalency of the applied phantom on the resulting Hounsfield look-up table (HLUT), HLUTs were specified based on each of the two phantoms. Here, three different approaches were used: 1) Hybrid approach described in the consensus guide including both the phantom inserts and the tabulated human tissues in the linear regression, 2) including only the tabulated human tissues in the linear regression, or 3) including only the phantom inserts in linear interpolation. For both 1) and 2), the piecewise linear curve was defined applying the strategy for the placement of the connection points described in the consensus guide. The resulting HLUTs are given by their connection points in the sheet "HLUTs". As in the consensus guide, connection point number 10 at 2000 HU is only stated to illustrate the slope of the HLUT in the bone region and is not an actual connection point. Also note that for the phantom-inserts-only HLUT, the number of connection points differ, as there are different numbers of phantom inserts for the two phantoms, and each insert corresponds to one connection point.
HLUT accuracy
To evaluate the HLUT accuracy, three additional datapoints were plotted. An additional cortical bone insert from the CIRS EPTN (tissue-equivalent) phantom was scanned with the same CT settings as for the two phantoms above. However, this phantom was scanned with a Siemens Somatom Drive CT scanner, which has been tested at commisioning and yearly QA to provide CT numbers for head-sized phantoms similar to the ones derived from the Siemens Somatom Confidence CT scanner. This insert consisted of cortical bone mimicking material in the full insert diameter of 3 cm. Please note, this phantom insert also comes in a variant where only the inner core of 1 cm consists of cortical bone mimicking material, while the outer part of the insert consists of solid water. However, this insert variant was not used here, since the consensus guide suggests to only use phantom inserts with a diameter of at least 2 cm to avoid partial volume effects.
The measured CT number is also given in the sheet "PhantomInserts" along with its batch-specific density and chemical composition and calculated SPR at 100 MeV. Moreover, two datapoints for biological bone materials were taken from the literature (Taasti et al., Phys Med Biol (2018), and Möhler et al. Phys Med Biol (2018)). These two biological bone materials were scanned with a Siemens Somatom AS and Flash CT scanner, respectively, which are equivalent to the Confidence and Drive scanner used in this investigation. Both biological bone materials were scanned at 120 kVp with similar reconstruction settings as used here. The SPR of the biological bone materials were experimentally measured.
The measured (CT number in HU, SPR) for the biological bone materials (see tables of original articles) were (1607.6 HU, 1.737) and (1559.4 HU, 1.7224).

## PhantomInserts
CIRS 062M
Insert name | Tissue group | Mean CT number (HU) | SPR at 100 MeV | Density | H | B | C | N | O | Na | Mg | Si | P | S | Cl | Ca | Ba
Lung (Inhale) | 1 | -793.4 | 0.17626 | 0.178 | 8.8 | 0 | 67.5 | 3.5 | 18.6 | 0 | 0 | 0 | 0 | 0 | 1.6 | 0 | 0
Lung Exhale | 1 | -523.27 | 0.48896 | 0.494 | 8.87 | 0 | 66 | 2.38 | 20.41 | 0 | 0 | 0 | 0 | 0 | 0.61 | 1.73 | 0
Adipose | 2 | -67.38 | 0.96817 | 0.962 | 9.97 | 0 | 71.36 | 1.79 | 16.37 | 0 | 0 | 0 | 0 | 0 | 0.19 | 0.32 | 0
Breast | 2 | -36.83 | 0.98403 | 0.983 | 9.6 | 0 | 70.33 | 1.92 | 17.02 | 0 | 0 | 0 | 0 | 0 | 0.19 | 0.94 | 0
Liquid water | 3 | -2.82 | 1 | 1 | 11.19 | 0 | 0 | 0 | 88.81 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Liver | 3 | 46.43 | 1.0536 | 1.062 | 8.97 | 0 | 69.45 | 2.12 | 17.14 | 0 | 0 | 0 | 0 | 0 | 0.13 | 2.19 | 0
Muscle | 3 | 51.64 | 1.058 | 1.065 | 9.07 | 0 | 69.78 | 2.07 | 16.8 | 0 | 0 | 0 | 0 | 0 | 0.12 | 2.16 | 0
Bone 200 | 4 | 237.32 | 1.1033 | 1.155 | 7 | 0 | 56.3 | 2 | 22.7 | 0 | 0 | 0 | 3.3 | 0 | 0.2 | 8.5 | 0
Bone 800 | 4 | 962.49 | 1.3991 | 1.52 | 5.67 | 0 | 40.81 | 1.01 | 25.88 | 0 | 0 | 0 | 8.32 | 0.07 | 0.03 | 17.93 | 0.28
Bone 1250 | 4 | 1461.35 | 1.6006 | 1.81 | 3.6 | 0 | 28.82 | 1.09 | 31.99 | 0 | 0 | 0 | 10.79 | 0.08 | 0.04 | 23.27 | 0.32
Gammex Advanced Electron Density phantom
Insert name | Tissue group | Mean CT number (HU) | SPR at 100 MeV | Density | H | B | C | N | O | Na | Mg | Si | P | S | Cl | Ca | Ba
Lung LN300 | 1 | -707.09 | 0.28027 | 0.291 | 7.43 | 0 | 57.86 | 1.96 | 20.71 | 0 | 11.19 | 0.77 | 0 | 0 | 0.08 | 0 | 0
Lung LN450 | 1 | -518.96 | 0.45377 | 0.471 | 7.44 | 0 | 58.03 | 1.97 | 20.69 | 0 | 11.22 | 0.57 | 0 | 0 | 0.08 | 0 | 0
HE General Adipose | 2 | -70.94 | 0.96468 | 0.962 | 9.73 | 0.05 | 71.41 | 2.71 | 14.35 | 0.18 | 0 | 1.11 | 0 | 0 | 0.12 | 0.34 | 0
HE Breast 50:50 | 2 | -40.53 | 0.98148 | 0.983 | 9.48 | 0.04 | 70.23 | 2.47 | 15.13 | 0.16 | 0.61 | 1.01 | 0 | 0 | 0.11 | 0.74 | 0
Liquid Water | 3 | 1.48 | 1 | 1 | 11.19 | 0 | 0 | 0 | 88.8106 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
HE CT Solid water | 3 | 1.04 | 1.0068 | 1.024 | 8.41 | 0.05 | 66.97 | 2.16 | 18.49 | 0.17 | 1.1 | 1.08 | 0 | 0 | 0.13 | 1.43 | 0
HE Brain | 3 | 32.98 | 1.0305 | 1.051 | 8.23 | 0.04 | 65.75 | 2.05 | 19.7 | 0.15 | 1.23 | 0.94 | 0 | 0 | 0.13 | 1.79 | 0
HE Liver | 3 | 61.47 | 1.061 | 1.081 | 8.25 | 0.03 | 66.87 | 2.25 | 19.02 | 0.1 | 0.75 | 0.65 | 0 | 0 | 0.14 | 1.94 | 0
HE Inner Bone | 4 | 305.56 | 1.146 | 1.21 | 6.38 | 0.03 | 53.79 | 1.73 | 25.64 | 0.11 | 1.68 | 0.72 | 0 | 0 | 0.1 | 9.82 | 0
CaCO3 30% | 4 | 474.67 | 1.2481 | 1.332 | 5.6 | 0 | 51.77 | 2.12 | 28.47 | 0 | 0.18 | 0 | 0 | 0 | 0.09 | 11.77 | 0
CaCO3 50% | 4 | 877.87 | 1.4147 | 1.559 | 4.03 | 0 | 40.34 | 1.52 | 34.13 | 0 | 0.3 | 0 | 0 | 0 | 0.07 | 19.62 | 0
HE Cortical Bone | 4 | 1489.45 | 1.6886 | 1.931 | 2.31 | 0 | 27.41 | 0.85 | 39.95 | 0 | 2.82 | 0 | 0 | 0 | 0.04 | 26.62 | 0
CIRS EPTN phantom - not used for calibration
Insert name | Tissue group | Mean CT number (HU) | SPR at 100 MeV | Density | H | B | C | N | O | Na | Mg | Si | P | S | Cl | Ca | Ba
Cortical Bone | 4 | 1433.72 | 1.6637 | 1.89 | 0.033 | 0 | 0.2537 | 0.0091 | 0.3528 | 0 | 0.0336 | 0 | 0.0882 | 0 | 0.0003 | 0.2291 | 0

## TabulatedHumanTissues
Estimated CT numbers (HU)
Tissue name | Tissue group | Density | Schneider; Gammex | Ödén; Gammex | Schneider; CIRS | Ödén; CIRS
Lung inflated | 1 | 0.26 | -742.1978 | -740.6327 | -742.2606 | -740.9323
Adipose tissue 3​ | 2 | 0.93 | -88.2518 | -88.9297 | -84.3837 | -87.7554
Adipose tissue 2 | 2 | 0.95 | -67.446 | -67.2413 | -64.0127 | -66.3317
Adipose tissue 1 | 2 | 0.97 | -46.6789 | -45.5809 | -43.6863 | -44.9391
Skeleton – yellow marrow​ | 2 | 0.98 | -38.8227 | -39.1372 | -34.9768 | -38.0286
Mammary gland​ 2 | 2 | 1.02 | 4.9837 | 8.37 | 6.7034 | 8.2605
Skeleton – red marrow​ | 2 | 1.03 | 8.752 | 11.2276 | 10.9202 | 11.3743
Adrenal gland​ | 2 | 1.03 | 16.3646 | 20.3476 | 17.6684 | 20.0048
Brain – Cerebrospinal Fluid | 3 | 1.01 | 12.3787 | 19.4851 | 11.7718 | 18.0857
Small intestine (wall)​ | 3 | 1.03 | 22.0922 | 27.995 | 22.3616 | 27.0623
Urinary bladder – filled | 3 | 1.03 | 28.0938 | 34.9881 | 27.4384 | 33.5684
Gallbladder – bile​ | 3 | 1.03 | 27.6657 | 34.2948 | 27.4105 | 33.0747
Pancreas​ | 3 | 1.04 | 29.5782 | 34.9673 | 30.1096 | 34.1791
Brain – white matter​ | 3 | 1.04 | 29.8197 | 35.0292 | 30.1397 | 34.1537
Prostate | 3 | 1.04 | 30.4116 | 36.6731 | 30.6358 | 35.6933
Testis | 3 | 1.04 | 32.3843 | 38.5839 | 32.4206 | 37.5196
Brain – grey matter | 3 | 1.04 | 34.0143 | 40.328 | 33.7533 | 39.1188
Stomach | 3 | 1.05 | 38.1804 | 43.9413 | 38.6272 | 43.0876
Muscle – skeletal​ 2 | 3 | 1.05 | 35.401 | 41.2015 | 35.6731 | 40.2675
Heart​ 2 | 3 | 1.05 | 38.3704 | 44.1844 | 38.6113 | 43.2328
Aorta | 3 | 1.05 | 41.4822 | 47.2732 | 40.7135 | 45.8612
Kidney​ 2 | 3 | 1.05 | 40.2589 | 46.1835 | 40.1698 | 45.0736
Ovary | 3 | 1.05 | 42.0526 | 48.4257 | 41.9654 | 47.2881
Liver 2 | 3 | 1.06 | 47.6998 | 53.723 | 47.6995 | 52.6422
Spleen | 3 | 1.06 | 48.7728 | 54.9963 | 48.7526 | 53.894
Blood – whole​ | 3 | 1.06 | 47.8697 | 54.1088 | 47.845 | 53.0048
Skeleton – cartilage​ | 4 | 1.1 | 100.2807 | 108.3543 | 97.3875 | 105.7877
Skeleton – spongiosa​ | 4 | 1.18 | 275.442 | 278.3028 | 253.9585 | 267.4415
Skeleton – cortical bone​ | 4 | 1.92 | 1471.2541 | 1481.0747 | 1353.926 | 1425.2129

## HLUTs
Hybrid approach of the consensus guide (including both phantom inserts and tabulated human tissues in the linear regression)
| Schneider; Gammex |  | Ödén; Gammex |  | Schneider; CIRS |  | Ödén; CIRS
Connection point | CT number (HU) | SPR | CT number (HU) | SPR | CT number (HU) | SPR | CT number (HU) | SPR
1 | -1024 | 0.0010656 | -1024 | 0.0010656 | -1024 | 0.0010656 | -1024 | 0.0010656
2 | -999 | 0.0010656 | -999 | 0.0010656 | -999 | 0.0010656 | -999 | 0.0010656
3 | -950 | 0.033199 | -950 | 0.034551 | -950 | 0.037925 | -950 | 0.038826
4 | -148 | 0.8494 | -149 | 0.84452 | -144 | 0.85539 | -148 | 0.84796
5 | -128 | 0.92279 | -129 | 0.92414 | -124 | 0.92366 | -128 | 0.92436
6 | -30 | 0.99816 | -33 | 0.99493 | -30 | 0.99657 | -30 | 0.99694
7 | 0 | 1 | 0 | 0.99519 | 0 | 1.0014 | 0 | 0.99728
8 | 71 | 1.0723 | 71 | 1.067 | 62 | 1.0643 | 64 | 1.0618
9 | 150 | 1.0908 | 158 | 1.0932 | 147 | 1.0845 | 156 | 1.0847
10 | 2000 | 1.9223 | 2000 | 1.9206 | 2000 | 1.886 | 2000 | 1.8729
HLUTs fitted only on tabulated human tissues
| Schneider; Gammex |  | Ödén; Gammex |  | Schneider; CIRS |  | Ödén; CIRS
Connection point | CT number (HU) | SPR | CT number (HU) | SPR | CT number (HU) | SPR | CT number (HU) | SPR
1 | -1024 | 0.0010656 | -1024 | 0.0010656 | -1024 | 0.0010656 | -1024 | 0.0010656
2 | -999 | 0.0010656 | -999 | 0.0010656 | -999 | 0.0010656 | -999 | 0.0010656
3 | -950 | 0.049302 | -950 | 0.04894 | -950 | 0.049357 | -950 | 0.049017
4 | -148 | 0.85265 | -149 | 0.84664 | -144 | 0.85669 | -148 | 0.84853
5 | -128 | 0.92596 | -129 | 0.92734 | -124 | 0.92506 | -128 | 0.92667
6 | -30 | 0.99966 | -36 | 0.99429 | -30 | 0.99754 | -34 | 0.99531
7 | 0 | 1.0009 | 0 | 0.99502 | 0 | 1.0009 | 0 | 0.99607
8 | 59 | 1.06 | 65 | 1.0598 | 59 | 1.06 | 64 | 1.0599
9 | 150 | 1.0938 | 158 | 1.0949 | 147 | 1.0967 | 156 | 1.0965
10 | 2000 | 1.9279 | 2000 | 1.923 | 2000 | 2.0071 | 2000 | 1.9584
HLUTs based on linear interpolation of the phantom inserts only
| Gammex |  | CIRS
Connection point | CT number (HU) | SPR | CT number (HU) | SPR
1 | -707.09 | 0.28027 | -793.4 | 0.17626
2 | -518.96 | 0.45377 | -523.27 | 0.48896
3 | -70.94 | 0.96468 | -67.38 | 0.96817
4 | -40.53 | 0.98148 | -36.83 | 0.98403
5 | 1.48 | 1 | -2.82 | 1
6 | 1.04 | 1.0068 | 46.43 | 1.0536
7 | 32.98 | 1.0305 | 51.64 | 1.058
8 | 61.47 | 1.061 | 237.32 | 1.1033
9 | 305.56 | 1.146 | 962.49 | 1.3991
10 | 474.67 | 1.2481 | 1461.35 | 1.6006
11 | 877.87 | 1.4147
12 | 1489.45 | 1.6886